SR1 Barrel 3 Shift CheckList

Version 1.1 04/04/05

Date/Time:                                       Shift Team:

Log Book Page:	First run number:

LMTs:	Last run number:

A. Preparing to Test Modules

- Update configuration files: Barrel3.xml and MURs.

- Check that xml files are ok

- cd ~/sct1/daqsoft/ and source setup_HEAD.csh
- cd sctConf and run ‘./checkSCTConfig’ and if wrong, edit files

- Generate DCS database for modules present in this run
- cd ~sct1/daqsoft/SctRodDaq, run ‘./copyDCS.sh’ to generate a map file
- at DCS PC, copy DCS database files from /afs/cern.ch/user/s/sctroddq/dcs’ to C://SCT_DCS/Configs/MoPs/
- reload DCS database by clicking ‘Load Mapping’
- select ‘BarrelView’ tab and check display matches that of sctGUI

- Ensure lights are OFF – check by looking at webcam

- (If enclosure has been opened, also check that both HEPA filters are OFF and that the air ioniser (at +Z end of barrel) is OFF

- Verify dry air flow into the enclosure. You will hear an audible noise from the pipe if it is running. Check that the pressure regulator is showing somewhere between 50 and 130 mbar.

- Start and/or check DCS environmental monitoring. Wait for the dew point to settle (typically -40C). Do not start run if above       -25C

- When the dew point has stabilized – fill in DCS check list

- The cooling expert should now switch on the cooling system. Wait for the stave temperature to stabilize at a value below 15C.

- Check LEDs are lit on all 4 interlock cards, once the temperature on the cooling pipes has dropped.

- Turn interlock key(s) to enable the SCT power supply system.

- Turn on the power to the PS crates

- Ensure BOC back panel is in place

B. Starting the System

Throughout the test sequences, the DCS conditions should be monitored REGULARLY (every 15 mins or so) – see Appendix I ‘Checking DCS conditions’ below.

- Select the “ELMB” panel

- The state of each ELMB must be toggling between 5 and 133

- If not, you need to reset the ELMBs by clicking on the ‘!’ right by CC NMT and clicking on Reset CC.

- If no DAQ running, restart DAQ by ‘cd /sct1/daqsoft/SctRodDaq’ and doing ‘./start’. Configure DAQ to view modules using DCS display of sctGUI.

- On Barrel View Mask ON LV channels, but DO NOT ramp up LV or HV

- Check – do you see both temperature sensors in sctGUI (check it is updating regularly), and do they correlate with the DCS panel?      YES/NO 
     - if no, DO NOT continue, the module connector is misplaced. Turn off crates etc and correct immediately!

- On Barrel View Mask on (ON) HV channels

- Check – are all HV channels flashing slowly?                 YES/NO

- if no, DO NOT continue, an HV channel may have failed. Mask off (OFF) immediately and call expert.
______________________________________________________
If this is the first test of the modules since oxford, do the following two steps:

- Switch on HV to 5V for each module. 
Are all channels ok (no trips)?                                          YES/NO
    - if no, DO NOT CONTINUE, an HV channel is shorted. Mask off immediately and call experts.

- Switch on HV to 50V for each module. 
Are all channels ok (no trips, and registering non-zero current less than 1uA)?                                                                      YES/NO
    - if no, DO NOT CONTINUE, an HV channel may have a problem. Mask off immediately and call experts.
______________________________________________________

- Ramp up LV and HV “On All Vdd in Steps” on both crates (database will be set for 150V)

- Do first DCS-CHECK (see Appendix 1)

- Resend module configurations (right mouse click on sctGUI)

- Probe modules after ABCD configured.

- Modules returning events on all expected RX channels? YES / NO

- At the PS console, set the “BarrelValues” panel to view detector leakage current.  Take a reference point.

- Do Hard Reset and re-Probe modules

- All modules returning CLK/2?   YES / NO

- At the PS console, re-read the leakage currents. Any which have changed by a factor of more than 2 will now be highlighted.

  - Open the BOC door and re-read the currents.  If any high readings now fall to normal values, there is a light leak!

C. Testing modules

It is foreseen to do two test sequences for barrel 3 modules. The first is a short test sequence to enable a comparison with Oxford data. A longer test sequence involves a full characterisation of the modules.

Do DCS-CHECK

- Return to the “initial” state, then “configured”, then “Start”.

-- Write down the Run number

- Do RXThresholdBasedOnConfigRegister test

- Adjust the thresholds in the MUR files if MAX < 200 or MIN > 60

- Return down to “initial” state (STOP, UNLOAD, SHUTDOWN)

- Come back to “start” state (BOOT, CONFIG, START)
_____________________________________________________________
1. Short test sequence for Oxford comparisons:

- Do NMASK test and check pass

- Do PipeLine test and check pass

- Do FullByPass test and check pass

- Do StrobeDelay test and check data

- If anything fails do comparison with Oxford SCTRODDAQ and/or SCTDAQ reference

- You should see both rising and lower edge of the top-hat

- Do ThreePointGain and check data

- Do full comparison with Oxford SCTRODDAQ and/or SCTDAQ reference data

- Check INPUT NOISE, OFFSET, GAIN, NOISE, RMS Offset etc

Do DCS-CHECK

- Do NoiseOccupancy and check data

- Do full comparison with Oxford SCTRODDAQ and/or SCTDAQ reference data

- Check Occupancy < 1E-4

- Do DoubleTriggerNoise and check data

- Average and peak should be < 1E-4

_____________________________________________________
2. Full Test Sequence:

- Do a full characterisation. Perform DCS check every 15 minutes.

- Cross check with previous SR1 data and/or Oxford SCTRODDAQ data

- Do DoubleTriggerNoise and check data

- Average and peak should be < 1E-4

______________________________________________________

- Check IV of sensors

- HV_Volt to 150 V

- allow a minute to settle

- check all HV_Curr < 1.5 uA

D. After Testing

- Ramp down LV and HV

- Mask off LV and HV channels

- Power off PS crates

- Turn and remove Interlock key to disable the SCT power supply system.

- Switch OFF cooling

- Wait for temperatures to rise well above dew point

Revision History

Version 1.0 by DR on 03/04/05. FIRST VERSION, derived from Oxford checklist.

Version 1.1 by DR on 04/04/05. Implement comments from Richard.

DCS Checks – responsibility of: ___________ (name)

Before cooling starts:

(Values in enclosure:)

Dew Point:						Humidity:

If dew point >9C then:

Ask mounting supervisor to turn on dry air

Wait until dew point falls <9C

Every 20 minutes:

| Time | T1,T0 < 35C | T1,T0 within 2C of neighbours | T¦ < 2C | HV_curr <1.5uA | Staves 10-15C | Dew Point |
| --- | --- | --- | --- | --- | --- | --- |

ACTIONS:

- If dew point is within 3C of stave
- turn off LV and HV, and mask off
- check/increase air flow

- If stave temperatures >15 C 
- turn off LV and HV
- mask off

- If T0 or T1 > 35 C			
- turn off LV and HV on at least some modules
- mask them off

If |T| > 3C

- check with Oxford data

During IV check:

If HV_Curr is >1.5 uA compare with previous

If HV_Curr is >2 uA. Turn off HV.

Any values outside the above limits should also be reported on the problem page.

Tested (LMT, Z+-, Upper/Lower)

Signed off:

Problems:

| s/n | LMT | “row” | Z+- | Problem | Informed |
| --- | --- | --- | --- | --- | --- |